## ET鸿鹄版本20200509
Bluetooth ear thermometer ET command frame format
NO | Types |  | start character (4Byte) |  |  |  | data length(2Byte) |  | function(1Byte) | date block(1,2,3...n) | checksum(1Byte) | end character(2Byte) |  | Notes
1 | temperature | MCU→APP | FA | AA | AA | AF | 00 | by calculation | 03H | 00(measuring position) 0000(measured value) 00(year) 00(month) 00(day) 00(hour) 00(minute) | by calculation | F5 | 5F | Variable data length
Introductions： 1、The measuring position is 1 Byte，position=01，indicating ear temperature；position=02，indicating forehead temperature；position=03, indicating object temperature； 2、The measured value is 2 bytes, bit15 is the unit marker bit, bit15=1, meaning Fahrenheit, bit15=0, meaning Celsius; 3、The time format is: year/month/day/hour/minute 1 byte each, for example: 2017/02/21 12:00, then send format: 17/02/21/12/00; 4、N memory values are sent continuously, starting at the first Byte and ending at the last. Each memory value is 8 bytes, N≤30, and the memory order is from new to old; 5、The measured value defaults to high precision(0.01 resolution) and Celsius. Please convert it according to the unit and convert it to 0.1 resolution (the unit must be converted before the precision is converted);
2 | time synchronization | MCU→APP | FA | AA | AA | AF | 00 | 0AH | 06H | / | 10H | F5 | 5F
Introductions： 1、As soon as bluetooth is connected, the MCU must immediately send a sync command to the BLE to keep the time of the measurement up to date;
Bluetooth ear thermometer ET response frame format
NO | Types |  | start character (4Byte) |  |  |  | data length(2Byte) |  | function (1Byte) | date(1,2,3...n) | checksum(1Byte) | end character(2Byte) |  | Notes
1 | temperature | App→MCU | FA | AA | AA | AF | 00 | 0AH | 83 | / | 8DH | F5 | 5F
2 | time synchronization | App→MCU | FA | AA | AA | AF | 00 | 0FH | 86 | 00(year) 00(month) 00(day) 00(hour) 00(minute) | by calculation | F5 | 5F
Note: The checksum is the low byte of the hexadecimal sum of all bytes of data length, function and data block;
Notes： | Bit15 unit flag bit, keep, but app does not respond
|  | for example: | FA AA AA AF 00 0F 86 14 06 13 0E 24 F4 F5 5F
FA AA AA AF 00 0A 83 8D F5 5F
After the connection is successful
The app receives the device's message through the UUID channel fff1
The app send message to the device through the UUID channel fff2
Our communication protocol is analyzes the data with these two channels to obtain useful information
Take the NRF connect software for example