The device is equipped with 12 sensing points, but data is transmitted for only 10 of them (starting from the sensor head). This is due to beacon capacity limitations and the assessment that 12 points are unnecessary.

Upon initial activation, the sensor undergoes a roughly 10-second initialization process to calibrate and save values based on the environment. Sensing is then performed using a combination of 10 initialized baseline values and 10 continuously changing values.

Sensing is achieved by measuring the degree of change—calculated as the difference between the baseline value and the current value.

Users can adjust the sensitivity settings. For example, a notification can be triggered if the number of sensors showing a change exceeds a specific threshold.

Thus, if a user sets the sensitivity parameters to a “range of 5” and a “threshold value of 10,” a notification is issued when more than five sensors exceed that threshold value of 10.

Although the sensor’s internal algorithm generates event data (such as “soiled” vs. “clean” status), we do not currently use these specific values for detection.

This is because desired sensitivity levels vary by company, and factors such as diaper type, temperature, and humidity also influence the readings.

Therefore, you only need to utilize the following data transmitted by the sensor: battery status, the 10 baseline values for the sensing points, and the 10 corresponding changing values.

(You can essentially disregard the notification values coming from the sensor.)

Information such as signal strength and MAC addresses is also included, so you can use it as needed.

I will provide detailed documentation at a later time.

Below is a concise draft covering Sections 3.3 through 5.3.

3.3 Packet Validation

A BLE advertisement is considered a valid MECSPro sensor packet when:

Manufacturer Specific Data is at least 23 bytes.

Byte 0 is 0x02.

Byte 1 is 0x15.

The following 20 bytes contain UUID (16 bytes), Major (2 bytes), and Minor (2 bytes).

No additional CRC or checksum validation is currently applied.

4. Raw20 Packet Layout

The 20-byte payload is created as follows:

Raw20 = UUID[16 bytes] + Major[2 bytes] + Minor[2 bytes]

The payload is parsed as one continuous 160-bit stream in MSB-first order.

Bit Offset	Length	Field	Range

0–2	3 bits	Packet Type	0–7

3–9	7 bits	Battery	0–127

10	1 bit	Alarm Type	0–1

11–12	2 bits	TX Strength	0–3

13–15	3 bits	Event Status	0–7

16–75	60 bits	Baseline Values	10 × 6 bits

76–135	60 bits	Raw Sensor Values	10 × 6 bits

136–159	24 bits	MAC Address Suffix	3 × 8 bits

There is no padding between fields. Some 6-bit sensor values cross byte boundaries, so the data must be parsed as a bit stream rather than as individual bytes.

4.1 Header Layout

The first two bytes contain the packet header.

Byte 0:

Bits 7–5: Packet Type

Bits 4–0: Battery bits 6–2

Byte 1:

Bits 7–6: Battery bits 1–0

Bit 5:    Alarm Type

Bits 4–3: TX Strength

Bits 2–0: Event Status

4.2 Byte Mapping

Bytes 0–1:   Header

Bytes 2–9:   Baseline values and beginning of sensor values

Bytes 9–16:  Remaining raw sensor values

Byte 17:     MAC suffix byte 1

Byte 18:     MAC suffix byte 2

Byte 19:     MAC suffix byte 3

Baseline and sensor values must be read sequentially in 6-bit units.

5. Field Interpretation

5.1 Sensor Values

The packet contains 10 sensor channels.

Each channel provides:

A 6-bit baseline value: 0–63

A 6-bit raw sensor value: 0–63

The normalized sensor value is calculated as:

Normalized[i] = max(RawSensor[i] – Baseline[i], 0)

The normalized values are used by the application for status evaluation and display.

5.2 Device MAC Address

The final three bytes of Raw20 contain the last three bytes of the device MAC address.

The application uses the following fixed prefix:

EE:C5:00

The full embedded MAC address is reconstructed as:

EE:C5:00:XX:YY:ZZ

Example:

MAC suffix: 12 34 56

Full MAC:    EE:C5:00:12:34:56

Note that the Android application currently identifies registered devices using the BLE address provided by the Android scan result. The relationship between this address and the embedded MAC address should be confirmed during integration testing.

5.3 Status Fields

Field	Description

Packet Type	3-bit packet classification value. Detailed code definitions must be confirmed with the firmware team.

Battery	7-bit battery value. The current application uses it directly as the battery level.

Alarm Type	1-bit alarm-type flag. The exact meaning of 0 and 1 must be confirmed.

TX Strength	2-bit transmission strength or sensitivity code.

Event Status	3-bit status value received from the sensor.

The application also calculates its own event status using the normalized sensor values:

Status 2:

All 10 normalized values are below 4.

Status 1:

The number of channels equal to or above the configured threshold

Is greater than or equal to the required channel count.

Status 0:

All other conditions.

Current default settings:

Required channel count: 4

Sensor threshold:       12

Sensitivity Settings

The sensitivity setting consists of two parameters:

PollutionRange: Number of sensor channels that must exceed the threshold. The valid range is 2–10.

PollutionValue: Threshold applied to each normalized sensor value. The valid range is 5–25.

Na alert is generated when:

Number of channels with NormalizedValue >= PollutionValue

Is greater than or equal to PollutionRange.

The application provides three recommended presets:

Preset	PollutionRange	PollutionValue

More Diaper Alerts	3	7

Normal Diaper Alerts	4	12

Fewer Diaper Alerts	7	15

Lower values increase sensitivity and generate more alerts. Higher values reduce sensitivity and generate fewer alerts.